|  |  |
| --- | --- |
| Name | Security Level |
| BELLSING® Hornet Series AMBA Receiver | Open |
| Version | A Total of 4 Pages |
| REV A |

BRC740C90025

Product Data Sheet

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Prepared by: | Sam.Chen |  | Date: | 2020-08-11 |
| Auditor: |  |  | Date: |  |
| Approver: |  |  | Date: |  |



Shenzhen Bellsing Acoustic Tech. Co., Ltd.

BRC740C90025 AMBA receiver

# Description

## Miniature balanced-armature receiver with best vibration performance for use in high end hearing aid instruments with great response. Also plays an excellent role as woofer in in-ear monitors and other commercial hearing products.


|  |  |
| --- | --- |
| Name: | Hornet Series AMBA Receiver |
| Part Number: | BRC740C90025 REV A |

# Construction Specification

## 

## Weight: 0.31 g

## Mechanical dimension：please refer to the drawing below



\*A positive voltage applied to the negative terminal (­) will result in an increase in pressure at the sound outlet.

# Features & Benefits

## Perfect package size for BTE, headset applications

## Outstanding efficiency

## Substantial maximum output

## Magnetic-radiation shielded

## Optimized mechanical shock resistance

# Electro-acoustic Specifications

## Test condition

## Acoustic tubing: 10.0mm of 1.0mm diameter tubing

## Acoutic coupler: IEC711 coupler

## Constant voltage drive: 0.110V RMS

## Environment: 23℃, RH 50%

## Acoustic Parameter I

|  |  |  |  |
| --- | --- | --- | --- |
| Test Item | Nominal | Tolerance | Comments |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Sensitivity(dB) | @20Hz | 115.5 | +/- 3.0 |  |
| @200Hz | 115 | +/- 3.0 |  |
| @500Hz | 112.5 | +/- 3.0 |  |
| @100Hz | 112.5 | +/- 3.0 |  |

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Peak1 | Frequency (Hz) | 1700 | +/- 350 |  |
| SPL(dB) | 120.5 | +/- 3.0 |  |
| Peak2 | Frequency (Hz) | 3750 | +/- 350 |  |
| SPL(dB) | 115.5 | +/- 3.0 |  |
| Maximum output@Peak frequency (dB) | | N/A | N/A | THD<10% |


## Acoustic Parameter II

|  |  |  |  |
| --- | --- | --- | --- |
| Test Item | | Nominal | Comments |
| THD(%) | @1/2 1st peak | <5.0 |  |
| @1/3 1st peak | <5.0 |  |

## Electric Parameter

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Test Item | | Nominal | Tolerance | Comments |
| Impedance(Ω) | @500Hz | 41.3 | +/- 15% |  |
| @1000Hz | 82.5 | +/- 15% |  |
| DCR(Ω) | | 25 | +/- 10% |  |

## Nominal Frequency Response Curve



# Package and Storage Condition

## Package: PET bubble pack, 1 pcs per bubble, 100 pcs per pack

## Storage temperature range: -40℃ ~ 60℃

## Storage humidity range: <60%